## Slide 1: Véronique Vallet, PhD Institute of Radiation Physics
Beam charaterisation

## Slide 2: Outline
Measurements means of dose in radiotherapy
Dosimetry parameters of photons beams
Use of the dosimetry parameters

## Slide 3: Measure of ionising beam
Indirect measurement (effects measurement)
Unit

## Slide 4: Ionising radiation
Electrons - direct:
  Continuously interacts with matter through Coulomb forces
  Energy loss in form of ionizations and excitations along the path
Photons - Indirect:
  “long” distance between successive interactions
  Interaction → often release of an electron
  Released electron loss its energy in form of ionizations and excitations

## Slide 5: Ionisation based measurement devices
Ionization of material (production of + and - charges)
Migration of charges to electrodes
Measurement of current (or charge) by electrometer
Electrometer signal increases with voltage
-
-
-
-
+
+
+
+
-
+
+
-
-
+
-
+
+
-
-
+
-
+
+
-
-
+
-
+
+
-
-
+
-
+
+
-
-
+
-
+
+
-
-
+
-
+
+
-
-
+
-
+
+
-
-
+
-
+
+
-
-
+
-
+
+
-
-
+
Ionizing radiation
Secondary electron
Positive ion
Negative ion
Negative electrode
Positive electrode
A
+
electrometer
HV

## Slide 6: Ionisation based measurement devices
Recombination before reaching the electrode
No recombination
Produced ionized particles produces other ionisations

## Slide 7: Chambre d’ionisation en radiothérapie
Voltage high enough to prevent recombination
Voltage low enough to prevent the created ions from ionizing
Cylindrical shape with central electrode
Sensitive volume size according to use
Sealed or ventilated
-
-
+
+
Incident photon
Scattered photon
e-
Central electrode
Sensitive volume

## Slide 8: Charge - Dose
: chamber calibration factor
: Energy correction factor (beam quality correction factor)
: pressure and temperature corerction factor ()
: recombination correction factor

## Slide 9: Reference condition
Field size
Fixe distance
  SAD = source – axe distance (isocentric)
  SSD = source – surface distance
Depth
target
Measurement point =
isocentre
Isocentric condition
SAD
target
measurement point
fixed SSD condition
SSD

## Slide 10: Machine calibration
«counts» measurement by the monitor chamber
Dose measurement by the reference ionisation chamber
In reference condition:
1 UM = 1 cGy
# countsm
Dm
# counts1UM
0.01 Gy
=
# counts1UM
# countsm
Dm
0.01
=
Monitor chamber calibration

## Slide 11: Field profil
Measured in a water tank with an ionization chamber

## Slide 12: Field profil
Parameters determining the quality of the measurement:
Detector size
Measurement interval

## Slide 13: Output Factor (FOC)
Field size effect:
ratio of the dose measured under reference conditions and the dose measured for a field size T, other conditions being similar
Dref
DT
=

## Slide 14: Output Factor
Head and phantom scatter
Detector size vs. field size → detector can disturb field
Use of diodes or diamond detector + Monte-Carlo calculated correction factors for output factor of small fields

## Slide 15: Percentage Depth Dose
Depth effect in fixed SSD condition:
Ratio of the dose measured under reference conditions and the dose measured for a depth p, other conditions being similar
Dref
Dp
=

## Slide 16: Percentage Depth Dose
Three zones:
Build-up
Electronic equilibrium
Exponential decay

## Slide 17: Photons → Indirectly ionising
Energy is not deposited where photons interacts with matter
Ahnesjö and Aspradakis, PMB 44 (1999), R99-155

## Slide 18: Interaction photon-matter
 |  | Quantité d’énergie cédée à la matière E |  | 
 |  | E = E0 (absorption) | 0 < E < E0 (inelastic diffusion) | E = 0 (elastic diffusion)
Type d’interaction | with electrons | Photoelectric | Compton | Thomson Rayleigh
 | with elm field | Pair creation |  | 
 | with nucleus | Photonuclear |  | 

## Slide 19: Photoelectric
Incident photon of energy
Photon absorption
Ejection of an electron of energy
Electron direction depends on the energy of the incident photon (the greater the energy, the greater the probability of emission in the same direction as the photon).

## Slide 20: Campton
Collision of an incident photon of energy with a free electron
Part of the energy is transferred to the electron ()
Remaining energy for the scattered photon
Energy-dependent scattering angles (the higher the energy of the scattered photon, the further forward it will travel)
Progressive energy loss with several successive Campton interactions

## Slide 21: Attenuation
Linear attenuation coefficient
=
Probability of at least one interaction per unit length for a photon in a beam
Exponential decrease of number of photon that did not interacted yet

## Slide 22: Percentage depth dose
Three zones:
Build-up:
  Electrons deposit energy
  Electrons move predominantly forward
  Lack of electrons on the surface due to absence of material before the surface
Electronic equilibrium
  As many electrons in as out
Exponential decay
  Less photons available to interact → less electrons produced

## Slide 23: Tissu phantom ratio (TPR)
Depth effect in isocentric condition Ratio of the dose measured under reference conditions andof the dose measured for a depth p, other conditions being similar
Dref
Dp
=

## Slide 24: Manual UM calculation
Retrieve reference conditions by applying correction factors
Fixed DSP conditions
Field size
OF
Depth
PDD
C
z
Cref
z
zref
Cref

## Slide 25: Manual UM calculation
Retrieve reference conditions by applying correction factors
Isocentric Conditions
Field size
OF
Depth
TPR
C
z
Cref
z
zref
Cref

## Slide 26: Manual UM calculation
Additional correction factors can be used to take account for exemple of:
Field shape (equivalent square)
Heterogeneity (equivalent depth)

## Slide 27: Beam modeling in treatment planning system
Reference conditions used directly
Profiles and PDD (or TPR) used to validate the model