Fundamentals in Biophotonics
Interaction of Light With Tissues-
Endoscopic techniques
Aleksandra Radenovic
aleksandra.radenovic@epfl.ch
EPFL – Ecole Polytechnique Federale de Lausanne
Bioengineering Institute IBI

---
Steps for development from Bench to Bedside

---
In 1967, a few years after the first working laser was invented, Endre Mester in Semmelweis University, Budapest, Hungary wanted to test if laser radiation might cause cancer in mice . He shaved the dorsal hair, divided them into two groups and gave a laser treatment with a low powered ruby laser (694 nm) to one group. They did not get cancer, and to his surprise the hair on the treated group grew back more quickly than the untreated group. This was the first demonstration of "laser biostimulation".

---
Interaction of Light With Tissues
A tissue is a self-supporting bulk medium light propagation produces
Absorption
Scattering
Refraction
Reflection
Plasma formation

---
Interaction of Light With Tissues
A tissue is a self-supporting bulk medium light propagation produces
Absorption
Scattering
Refraction
Reflection
Plasma formation

---
Light Scattering Processes
Most pronounced effect in a tissue
Turbidity – measurement of submicron particles that obscure light rays, or apparent nontransparency of a tissue results in a loss of the initial directionality of a collimated beam as well as in defocusing of the light beam spot

---
Optical Penetration depth
Scattering creates a loss of intensity as the light propagates through a tissue
Described by an exponential function
Scattering occurs as photons of light interact with and pass through tissue structures.
Most information obtained from studying the bulk scattering component of reflectance represents information about the presence of large particles, such as the presence of collagen in the lamina propria and submucosa

---
Optical Penetration depth δ
Measurement of how deep light can penetrate into a tissue
Initial intensity is reduced by 90% at a depth of 2δ
δ decreases with the vascularity (blood content) of a tissue
δ is less for blue light than for red light
Largest in the region of 800 – 1300 nm

---
Measurement of Optical Properties of a Tissue
Typical transmission experiment – measure transmission of a beam through a tissue of a finite length
Gives us a total attenuation coefficient
Double-Integrating spheres
Simultaneously determines the reflectance, absorption, and scattering
Measures the intensity of light in the forward direction of propagation
Two nearly identical spheres are located in front of and behind the tissue sample
Light detectors collect light from all angles (hence the term integrating spheres)

---
Absorption of Light in Tissue
The tissue molecules that absorb the light are usually referred to as pigments. Hemoglobin and water are two common body constituents that can function as pigments. Hemoglobin has a very high absorption in the violet and blue/green portions of the visible spectrum. Absorption declines in the red region of the spectrum, which is why hemoglobin is red - it does not absorb red light.
The graphic shows absorption spectra of major intracellular absorbers. The molecular exctinction coefficients of oxygenated haemoglobin and melanin and the absorption coefficient of water are shown.

---
Double-Integrating Sphere

---
Light-Induced Processes

---
Autofluorescence-tissue
Any given tissue has a nonuniform distribution of may fluorophores
The amount of fluorophores may vary as a function of depth below the tissue surface
Autofluorescence may be different from a premalignant or malignant tissue compared to a normal tissue
Development of cancer involves a series of changes some of which can be probed by fluorescence
protein expression (Trp)
metabolic activity (NADH/FAD)
nuclear morphology
organization
structural integrity (collagen)
angiogenesis

---
Instrumentation for clinical tissue fluorescence measurements
can be very simple, compact and relatively cheap
Light Detector
CCD camera
photodiode, phototube
streak camera
single frame, video
time gated
 wavelength selector
Image Acquisition
 wide-field
point measurements+scanning
 endoscopic
 proximity or remote
Light Source
lasers, LEDs, lamps
wavelength selectors
cw, oscillatory, pulsed

---
An imaging technique that uses the interactions of light with cells and tissues to diagnose and treat abnormalities.
Ex. When light shines on cells, the particles of light are scattered by atoms in the molecules of the cells. A special imaging device records these scatter patterns. The molecules in abnormal cells create different scatter patterns than normal cells
Pros: view tissues deep in the body using light
Endoscope: a thin flexible tube that has bright light and video cameras a its end.
Surgical techniques: endoscope has surgical tools at its end so it can be used with minimal small incisions and the doctor can perform any surguries by using an endoscope such as repair knees, gallbladder removal and more

---
Neoplasia –pancreatic cells
. Neoplasia ("new growth" in Greek) is the abnormal proliferation of cells.

---
Consistent autofluorescence differences have been detected
between normal, pre-cancerous and cancerous spectra
Non-dysplastic Barrett’s
esophagus
Low-grade dysplasia
High-grade dysplasia
Promising studies in
GI tract
Cervix
Lung
Oral cavity
Breast
Artery
Bladder
Wavelength (nm)
Normalized fluorescence intensity

---
Methods of data analysis
Main goal for fluorescence diagnostics: Identify fluorescence features that can be used to identify/classify tissue as normal or diseased.
Main approaches
  - Statistical
  - Empirical
  - Model Based
Empirical and statistical algorithms
Data pre-
processing
Data reduction
and
Feature extraction
Classification
Normalization
Principal Component
Analysis
Ratio methods

---
Detection of cervical pre-cancerous lesions using fluorescence spectroscopy
During the natural lifetime of a woman, squamous epithelium which lines the ectocervix gradually replaces the columnar epithelium of the endocervix, within an area known as the transformation zone. The replacement of columnar epithelium by squamous epithelium is known as squamous metaplasia.
Most pre-cancerous lesions of the cervix develop within the transformation zone.
The Papanicolaou (Pap) smear is the standard screening test for cervical abnormalities
If a Pap smear yields atypical results, the patient undergoes a colposcopy, i.e. magnified (typically 6X to 15X) visualization of the cervix.
3-6% acetic acid is applied to the cervix and abnormal areas are biopsied and evaluated histo
4-6 billion dollars are spent annually in the US alone for colposcopic evaluation and treatment
Major disadvantage colposcopic evaluation is its wide range of sensitivity (87-99%) and specificity (23-87%), even in expert hands.

---
Detection of cervical pre-cancerous lesions using fluorescence spectroscopy
ectocervix
ectocervix
endocervix
Colposcopic view of
uterine cervix
Transformation
zone
endocervix
Major tissue histopathological classifications
Normal squamous epithelium
Squamous metaplasia
Low-grade squamous intraepithelial lesion
High-grade squamous intraepithelial lesion
Carcinoma

---
Instrumentation

---
Data processing
        337 nm Excitation 380 nm Excitation 460 nm Excitation
PRE-PROCESSING
Normalized Spectra at
Three Excitation Wavelengths
Normalized, Mean-scaled Spectra
 at Three Excitation Wavelengths
DIMENSION REDUCTION: PRINCIPAL COMPONENT ANALYSIS
SELECTION OF DIAGNOSTIC PRINCIPAL COMPONENTS: T-TEST
CLASSIFICATION: LOGISTIC DISCRIMINATION
Constituent Algorithm 1
Constituent Algorithm 3
Constituent Algorithm 2
Posterior Probability of being NS or SIL
Posterior Probability of being LG or HG
Posterior Probability of being NC or SIL
DEVELOPMENT OF COMPOSITE ALGORITHMS
Composite Screening Algorithm
Composite Diagnostic Algorithm
(1,2)
(1,,3)

---
Data processing
Data
Normal squamous
Low-grade
High-grade
Normal columnar
Pre-
Processing
Step 1
Pre-
Processing
Step 2

---
Detection of lung carcinoma in situ using the LIFE imaging system
an area of abnormal brownish red fluorescence in the sub-carina. The area (image #2) measured 2 to 3 millimeters in width and was confirmed by biopsy as carcinoma-in-situ with small foci of microinvasion.
Carcinoma in situ
Autofluorescence ratio
image
White light bronchoscopy
[Speaker Notes: 68 year old male- Ex-smoker Presented with a persistent cough and sputum production Sputum cytology showed cells suggestive of squamous carcinoma The referring pulmonologist completed a white light bronchoscopy and was unable to localize the source of the malignant cells identified by sputum cytology. The patient was referred to another pulmonologist for examination with the Xillix LIFE-Lung Fluorescence Endoscopy System. A white light bronchoscopy was performed, followed by a fluorescence bronchoscopy using the Xillix LIFE-Lung System. In the right lower lobe it was noted that some thickening had occurred in bronchus RB8 (image #1), however, a similar thickening was also noted elsewhere in the bronchial tree such as in adjacent bronchus RB9. Frequently, endoscopists see patients like this but may or may not be able to identify the tissue for biopsy because the changes are subtle and cannot be readily distinguished from inflammation or benign metaplasia under conventional white light examination. When the Xillix LIFE-Lung System was used and RB8 was re-examined there was an area of abnormal brownish red fluorescence in the sub-carina. The area (image #2) measured 2 to 3 millimeters in width and was confirmed by biopsy as carcinoma-in-situ with small foci of microinvasion. No abnormal fluorescence was observed from the thickened adjoining bronchus RB9. This case demonstrates a situation where the Xillix LIFE-Lung System differentiated an area of carcinoma in-situ from a normal appearing sub-carina. This information could be vital in identifying the extent of the endo-bronchial spread and for determining the appropriate treatment modality.]

---
Fluorescence imaging based on ratio methods
Wide field of view (probably a huge advantage for most clinical settings)
Eliminates effects of distance and angle of illumination
Easy to implement
Provides no intuition with regards to origins of spectral differences
Collagen and NADH spectra are sufficiently distinct only for some excitation wavelengths
337 nm excitation
358 nm excitation
Wavelength (nm)
Wavelength (nm)

---
Tissue absorption and scattering may affect significantly tissue fluorescence
single scattering
epithelium
absorption
Hemoglobin, beta carotene
fluorescence
Connective
tissue

---
Fluorescence and Fluorescence Lifetime Imaging
Pancreas
Prostate
Bladder
FLIM and H&E images of melanoma cells
633 nm excitation NIR autofluorescence
Images showing cancer (CA) lesions.

---
Fundamentals in Biophotonics
Optical coherence tomography
Aleksandra Radenovic
aleksandra.radenovic@epfl.ch
EPFL – Ecole Polytechnique Federale de Lausanne
Bioengineering Institute IBI

---
Optical coherence tomography
The in situ imaging of tissue microstructure with a resolution approaching that of histology, but without the need for tissue excision and processing

---
OCT vs. standard imaging
Standard
clinical
High
frequency
OCT
1 mm
1 cm
10 cm
Penetration depth (log)
1 mm
10 mm
100 mm
1 mm
Resolution (log)
Confocal
microscopy
Ultrasound

---
Optical coherence tomography
Three-dimensional imaging technique with ultrahigh spatial resolution even in highly scattering media
Based on measurements of the reflected light from tissue discontinuities
  - e.g. the epidermis-dermis junction.
Based on interferometry
  - involves interference between the reflected light and the reference beam
Ophthalmology
diagnosing retinal diseases.
Dermatology
skin diseases,
early detection of skin cancers.
Cardio-vascular diseases
vulnerable plaque detection.
Endoscopy (fiber-optic devices)
gastrology,
…
Functional imaging
Doppler OCT,
spectroscopic OCT,
optical properties,
PS-OCT.
Guided surgery
delicate procedures
brain surgery,
knee surgery,
…

---
OCT setup
Fiber-optic
beamsplitter
Broadband
source
Tissue
Scanning
reference mirror
Detector
Amplifier
Bandpass filter
Computer

---
Optical coherence tomography
Challenge: scattering of photons destroys localization

---
System perspective
Light sources
Superluminescent diodes
Semiconductor amplifiers
Femtosecond lasers
…
Beam delivery and probes
Hand-held probe
Catheter
Ophthalmoscope
Microscope
OCT imaging engine
Resolution
Reference delay scanning
Doppler/polarization/spectroscopy
Detection
Frequency domain
continue
Computer control
Drive system
Real-time display
Data management
Image & signal processing
Motion reduction
Speckle reduction
Image enhancement
Rendering algorithms
…
DTU course (10380) - 2004

---
Choosing the light source
Four primary considerations
  - wavelength,
  - bandwidth,
  - power (in a single-transverse-mode),
  - stability;
    1. portability, ease-of-use, etc.
Light propagation (Monte Carlo simulation)
Absorption
“Snake” component
Incident light
Ballistic component
Diffuse reflectance
Diffuse transmittance

---
Choosing the light source
Scattering mainly attenuates the signal
Basic property
  - the temporal coherence envelope function G(t) is related to the power spectral function S(n) through
    G(t) = FT{S(n)}
    - Wiener-Kinchine theorem
  - broadband source  high axial resolution

---
Source spectrum and envelope
The axial resolution is
notice that Dl is the 3dB-bandwidth!

---
Light sources for OCT
Continuous sources
  - SLD/LED/superfluorescent fibers,
  - center wavelength;
    1. 800 nm (SLD),
    1. 1300 nm (SLD, LED),
    1. 1550 nm, (LED, fiber),
    1. power: 1 to 10 mW (c.w.) is sufficient,
  - coherence length;
    1. 10 to 15 m (typically),
Pulsed lasers
  - mode-locked Ti:Al2O3 (800 nm),
  - 3 micron axial resolution (or less).
Scanning sources
  - tune narrow-width wavelength over entire spectrum,
  - resolution similar to other sources,
  - advantage that reference arm is not scanned,
  - advantage that fast scanning is feasible.

---
OCT spatial resolution
Axial and lateral resolutions are decoupled
The lateral resolution is determined by the focusing conditions
  - optics
  - dynamic vs. static focusing
Lateral resolution
b
Depth of focus
Dz
Dz
Dz
High NA
Low NA
b
Dz

---
High resolution OCT
Broad bandwidth sources
  - solid-state lasers,
  - sub-5 fs pulse;
    1. Ti:Al2O3 (Spectral bandwidth: 350 nm demonstrated),
  - other lasers/wavelengths available or needed.
Special interferometers and fiber optics
  - support for broad spectral range,
  - dispersion balanced,
  - current system used for OCT: 260 nm bandwidth, ~1.5µm resolution.
Chromatically corrected optics
  - aberrations can decrease resolution and SNR.
Broad bandwidth detectors and electronics
  - dual balance detection,
  - low noise circuitry necessary

---
Scanning devices
Piezo or motorized scanning devices
  - ideal for both longitudinal and lateral scanning.
Galvanic mirrors
Resonance scanners
Helical mirrors
  - longitudinal scanning.
Fiber stretcher
  - longitudinal scanning

---
Fourier domain rapid scanning optical delay line (RSOD)
The technique was originally developed for femtosecond pulse measurements
  - based on Fourier-transform pulse shaping techniques.
Relies on the basic property of the Fourier transform
  - phase ramp in
  - the Fourier domain
  - corresponds to a group delay in the
  1. time domain.

---
RSOD in the lab
Free-space group pathlength
  - p: the grating pitch,
  - f: focal length,
  - s: mirror angle.
Interferogram central frequency
Bandwidth

---
RSOD
Scanning capabilities (galvo)
  -  200 Hz,
  -  5 mm.
Scanning capabilities (resonant)
  -  4-8 kHz,
  -  5 mm.
Advantages
  - dispersion compensation feasible since phase and group delays may be separated,
  - center frequency of interferogram is adjusted through axial position

---
OCT in non-invasive diagnostics
Functional imaging
Doppler OCT,
spectroscopic OCT,
optical properties,
PS-OCT.
Ophthalmology
diagnosing retinal diseases.
Dermatology
skin diseases,
early detection of skin cancers.
Cardio-vascular diseases
vulnerable plaque detection.
Endoscopy (fiber-optic devices)
gastrology,
…
Guided surgery
delicate procedures
brain surgery,
knee surgery,
…

---
OCT data volume showing the morphology of a human finger pad
Scanning depth is 2 mm.

---
OCT examples
High magnification OCT image of benign human bladder taken with a clinical system during cystoscopy. Three distinct layers are visible: urothelium U, lamina propria L, and muscularis propria M. A sharp border is visible between the urothelium and the lamina propria.
Low magnification OCT image of a neoplastic rat bladder taken with aresearch system. The diseased urothelium U’, arrow is clearly visible.
.

---
OCT imaging of the human breast

---
Intravascular imaging of a fibrous coronary plaque showing: i—intimia with intimial hyperplasia, m—media, and a—adventitia.
The left figure is an intravascular OCT image showing a clear delineation of layers, including the internal *.

---
Clinically adapted systems