Fundamentals in Biophotonics
Optical coherence tomography
Aleksandra Radenovic
aleksandra.radenovic@epfl.ch
EPFL – Ecole Polytechnique Federale de Lausanne
Bioengineering Institute IBI

---
Optical coherence tomography
The in situ imaging of tissue microstructure with a resolution approaching that of histology, but without the need for tissue excision and processing

---
OCT vs. standard imaging
Standard
clinical
High
frequency
OCT
1 mm
1 cm
10 cm
Penetration depth (log)
1 mm
10 mm
100 mm
1 mm
Resolution (log)
Confocal
microscopy
Ultrasound

---
Optical coherence tomography
Three-dimensional imaging technique with ultrahigh spatial resolution even in highly scattering media
Based on measurements of the reflected light from tissue discontinuities
  - e.g. the epidermis-dermis junction.
Based on interferometry
  - involves interference between the reflected light and the reference beam
Ophthalmology
diagnosing retinal diseases.
Dermatology
skin diseases,
early detection of skin cancers.
Cardio-vascular diseases
vulnerable plaque detection.
Endoscopy (fiber-optic devices)
gastrology,
…
Functional imaging
Doppler OCT,
spectroscopic OCT,
optical properties,
PS-OCT.
Guided surgery
delicate procedures
brain surgery,
knee surgery,
…

---
OCT setup
Fiber-optic
beamsplitter
Broadband
source
Tissue
Scanning
reference mirror
Detector
Amplifier
Bandpass filter
Computer

---
Optical coherence tomography
Challenge: scattering of photons destroys localization

---
System perspective
Light sources
Superluminescent diodes
Semiconductor amplifiers
Femtosecond lasers
…
Beam delivery and probes
Hand-held probe
Catheter
Ophthalmoscope
Microscope
OCT imaging engine
Resolution
Reference delay scanning
Doppler/polarization/spectroscopy
Detection
Frequency domain
continue
Computer control
Drive system
Real-time display
Data management
Image & signal processing
Motion reduction
Speckle reduction
Image enhancement
Rendering algorithms
…
DTU course (10380) - 2004

---
Choosing the light source
Four primary considerations
  - wavelength,
  - bandwidth,
  - power (in a single-transverse-mode),
  - stability;
    1. portability, ease-of-use, etc.
Light propagation (Monte Carlo simulation)
Absorption
“Snake” component
Incident light
Ballistic component
Diffuse reflectance
Diffuse transmittance

---
Choosing the light source
Scattering mainly attenuates the signal
Basic property
  - the temporal coherence envelope function G(t) is related to the power spectral function S(n) through
    G(t) = FT{S(n)}
    - Wiener-Kinchine theorem
  - broadband source  high axial resolution

---
Source spectrum and envelope
The axial resolution is
notice that Dl is the 3dB-bandwidth!

---
Light sources for OCT
Continuous sources
  - SLD/LED/superfluorescent fibers,
  - center wavelength;
    1. 800 nm (SLD),
    1. 1300 nm (SLD, LED),
    1. 1550 nm, (LED, fiber),
    1. power: 1 to 10 mW (c.w.) is sufficient,
  - coherence length;
    1. 10 to 15 m (typically),
Pulsed lasers
  - mode-locked Ti:Al2O3 (800 nm),
  - 3 micron axial resolution (or less).
Scanning sources
  - tune narrow-width wavelength over entire spectrum,
  - resolution similar to other sources,
  - advantage that reference arm is not scanned,
  - advantage that fast scanning is feasible.

---
OCT spatial resolution
Axial and lateral resolutions are decoupled
The lateral resolution is determined by the focusing conditions
  - optics
  - dynamic vs. static focusing
Lateral resolution
b
Depth of focus
Dz
Dz
Dz
High NA
Low NA
b
Dz

---
High resolution OCT
Broad bandwidth sources
  - solid-state lasers,
  - sub-5 fs pulse;
    1. Ti:Al2O3 (Spectral bandwidth: 350 nm demonstrated),
  - other lasers/wavelengths available or needed.
Special interferometers and fiber optics
  - support for broad spectral range,
  - dispersion balanced,
  - current system used for OCT: 260 nm bandwidth, ~1.5µm resolution.
Chromatically corrected optics
  - aberrations can decrease resolution and SNR.
Broad bandwidth detectors and electronics
  - dual balance detection,
  - low noise circuitry necessary

---
Scanning devices
Piezo or motorized scanning devices
  - ideal for both longitudinal and lateral scanning.
Galvanic mirrors
Resonance scanners
Helical mirrors
  - longitudinal scanning.
Fiber stretcher
  - longitudinal scanning

---
Fourier domain rapid scanning optical delay line (RSOD)
The technique was originally developed for femtosecond pulse measurements
  - based on Fourier-transform pulse shaping techniques.
Relies on the basic property of the Fourier transform
  - phase ramp in
  - the Fourier domain
  - corresponds to a group delay in the
  1. time domain.

---
RSOD in the lab
Free-space group pathlength
  - p: the grating pitch,
  - f: focal length,
  - s: mirror angle.
Interferogram central frequency
Bandwidth

---
RSOD
Scanning capabilities (galvo)
  -  200 Hz,
  -  5 mm.
Scanning capabilities (resonant)
  -  4-8 kHz,
  -  5 mm.
Advantages
  - dispersion compensation feasible since phase and group delays may be separated,
  - center frequency of interferogram is adjusted through axial position

---
OCT in non-invasive diagnostics
Functional imaging
Doppler OCT,
spectroscopic OCT,
optical properties,
PS-OCT.
Ophthalmology
diagnosing retinal diseases.
Dermatology
skin diseases,
early detection of skin cancers.
Cardio-vascular diseases
vulnerable plaque detection.
Endoscopy (fiber-optic devices)
gastrology,
…
Guided surgery
delicate procedures
brain surgery,
knee surgery,
…

---
OCT data volume showing the morphology of a human finger pad
Scanning depth is 2 mm.

---
OCT examples
High magnification OCT image of benign human bladder taken with a clinical system during cystoscopy. Three distinct layers are visible: urothelium U, lamina propria L, and muscularis propria M. A sharp border is visible between the urothelium and the lamina propria.
Low magnification OCT image of a neoplastic rat bladder taken with aresearch system. The diseased urothelium U’, arrow is clearly visible.
.

---
OCT imaging of the human breast

---
Intravascular imaging of a fibrous coronary plaque showing: i—intimia with intimial hyperplasia, m—media, and a—adventitia.
The left figure is an intravascular OCT image showing a clear delineation of layers, including the internal *.

---
Clinically adapted systems