4-point measurement technique

At this station, we want to measure small values of resistance precisely. To this aim we will use tyo methods: 2-points measurement and 4-points measurement.

2-point measurement
1. Use a circuit board and the source-meter (HP34401A) in Ohm-meter mode to measure the resistance of a wire. Note down the value.
2. You can use a heat-gun to vary the temperature in the various cables. What are your thoughts on this resistance measurement?

4-point measurement
1. Connect the wire in order to realize a 4-wire measurement (figure below, and see the HP34401Q manual online). How does this new measurement compare with the previous one?
2. Use the heat-gun and compare the effect with the previous observation.

Lockin amplification for small resistance measurement

A lock-in amplifier is an electronic instrument used to measure very small AC signals, even when those signals are buried in a lot of noise. It relies on the orthogonality of sine functions: if you multiply two sine functions of different frequencies and integrate over a long range, the result is close to zero. In a lockin amplifier, the signal is multiplied by a reference sine wave of adjustable frequency, and integrated over an adjustable time constant/fed in a low-pass filter. The result is significant only if the input a signal has a frequency component equal to that of the reference sine wave (figure below).

The goal of this practical workshop is to use a lockin amplifier to measure dynamical impedances in a 4-points configuration. For that, we will induce an oscillating current of adjustable frequency in the resistance, and measure the voltage drop across it with the lockin.

Lockin amplification
1. Download the SRS830 manual. Take a look at the “THE BASIC LOCK-IN” section and the “X, Y, R and Theta” section.
2. What are the X and Y channels? What does the frequency and phase of the lockin correspond to?
3. What is the “auto-phase” functionality doing?

Installation of the lockin setup for resistance measurement

1. The lockin has a voltage output (sine out). Use it in combination with a resistor to build a current source of adjustable frequency.
2. We will measure the resistance of a wire (typically in the range of a few mOhms). Install the wire in a “4-points” measurement configuration. Connect the two ends of the wire to the lockin input.

1. Set the lockin amplifier to a frequency of 5 Hz. What is the impedance of the wire? Does the measured value of resistance matches that of a DC 4-points measurement? What can you say stability-wise?
2. Increase the lockin frequency. What is the effect on the signal, both in amplitude and phase? Compare the results with the impedence of the big resistance used in the current source.
3. Make a plot of these quantities versus frequency. Can you give an interpretation of the result? The observed effect is known as the skin effect. It must be taken into account when designing cables that carry high-frequency signals. For instance, Litz wires, which consist of multiple thin strands insulated electrically from one another, are designed to carry alternating currents above 1 MHz.

Side-quest : Frequency response of a LC circuit

  1. Replace the resistance with a capacitor and an inductance of 22mH.
  2. Plot the amplitude and phase response of the system with respect to frequency, and fit it with the theoretical response of a RLC circuit. What is the capacitance of the capacitor you used?