## Slide 1: Radiography & Mammography
Dr. Sc. Damien RACINE
Institute of Radiation Physics (IRA), CHUV
Damien.Racine@chuv.ch

## Slide 2: Contents of the lecture today
Introduction (15 min)
  History of x-ray imaging
  Principles of x-ray imaging
Radiography (1h 15 min)
  Chain of image formation
  X-ray production
  X-ray beam
  Interaction with patient
  Detection systems
  Dosimetry
Image Quality (45min )
  Physical measures of image quality
  Acquisition parameters
  Influence of acquisition parameters on patient dose and image quality
Mammography

## Slide 3: Contents of the lecture today
Introduction
  History of x-ray imaging
  Principles of x-ray imaging
Radiography
  Chain of image formation
  X-ray production
  X-ray beam
  Interaction with patient
  Detection systems
  Dosimetry
Image Quality
  Physical measures of image quality
  Acquisition parameters
  Influence of acquisition parameters on patient dose and image quality

## Slide 4: Interaction of radiation with matter
Photoelectric effect at low photon energies (<26 keV) with a probability dependant on photon energy and the atomic number of absorber
Compton scattering dominates at higher energies with low Z materials.
Rayleigh scattering: low probability in medical imaging (10% of interactions in mammography, 5% in chest radiography).
Pair production: beyond the range of diagnostic and nuclear radiology.

## Slide 5: The dose in radiologie
Mass element
Deposit of energy [J]
Absorbed radiation Dose
  Physical quantity
  Absorbed Dose : D (gray : Gy)
  Energy per unit mass (J/kg)
DT,R
DT,R
HT
Weighting factors for radiation
WR
  Quantity of biological effects
  Equivalent Dose HT = D x wR (sievert : Sv)
∑
WT
Weighting factors for tissues
  Quantity taking into account radiosensitivity of organs
  Effective Dose E = S HTwT (sievert : Sv)
E
DR
One measure point
QR
Quality factors
Measured Dose
Local doses
Personnel doses
Dose indices
Estimation
Notes: Facteurs de pondération wR: photons et électrons = 1
Neutrons à basse énergie < 10 keV = 5
Neutrons à haute energie jusqu’à 2 MeV = 20
Protons = 5
Particules alpha = 20
Facteurs de pondération wT:
Moelle osseuse, côlon, poumon, estomac, sein, tissus restants* = 0.12
Gonades = 0.08
Vessie, œsophage, thyroïde = 0.04
Surface de l’os, cerveau, glandes salivaires, peau = 0.01

## Slide 6: Dosimetry of radiography systems: Kerma
Kerma: kinetic energy released in matter.
Unit: Joule per kilogram [J/kg]  special name: Gray
1 J kg-1 = 1 Gy
radiological constant of air kerma
Filtration
Distance to focus
Charge (mAs)
Air Kerma at 100 kV for 1 mAs at 1 m
Generic value :
0.1 mGy m2/mAs

## Slide 7: Dosimetry : Entrance dose
Entrance dose
Represents deterministic effects
C = 0.1
BSF = 1.35 (backscatter factor)
FKD = 1 (Kerma-dose conversion factor)

## Slide 8: Chest radiography – Exercise
What is the dose to the skin ?
(120 kV / 0.5 mAs / 1.8 m /
3mm filtration)

## Slide 9: Chest radiography – Exercise (solution)
What is the dose to the skin ?
(120 kV / 0.5 mAs / 1.8 m /
3mm filtration)

## Slide 10: Abdomen radiography – Exercise
What is the dose to the skin ?
(80 kV / 80 mAs / 1 m /
3mm filtration)

## Slide 11: Abdomen radiography – Exercise (solution)
What is the dose to the skin ?
(80 kV / 80 mAs / 1 m /
3mm filtration)

## Slide 12: Comparison
skin dose: 0.03 mGy
skin dose: 6.9 mGy
In both cases, we have 2 μGy to the detector
thickness
~18 cm
(density 0.3)
thickness
~18 cm
(density 1)

## Slide 13: Dosimetry : Dose Area Product
Dose Area Product (DAP)
DAP = K x S
K : Kerma
S : Area exposed
Invariant with the distance between source and patient
Represents stochastic effects

## Slide 14: Dosimetry : Effective dose
The effective dose (E) measures the deleterious impact of the absorbed dose on biological tissues.
Each tissue does not have the same radiosensitivity
  It is the tissue-weighted sum of the equivalent doses in all specified tissues and organs of the human body
  is the tissue weighting factor
  is the radiation weighting factor
  is the mass-averaged absorbed dose in tissue T
  is the equivalent dose absorbed by tissue T

## Slide 15: Dosimetry : Effective dose
The sum of these weighting factors is 1, which corresponds to the overall detriment of irradiation of the whole body.
The higher WT , the more sensitive the tissues or organs.

## Slide 16: Dosimetry : Effective dose
Limitation:
  The calculation of the effective dose requires numerical simulations or multiple measures which are costly in terms of calculation and time
For a simple estimation, the effective dose is calculated by multiplying the dosimetric quantity (DoQ) by a conversion factor (eDoQ)
Radiography
Conversion factor eDAP [mSv/(Gy∙cm2)]

## Slide 17: Exercise
What is the effective dose ?
Data 2
Occupational and patient exposure as well as image quality for full spine examinations with the EOS imaging system; Damet et al.
thorax

## Slide 18: Solution
What is the effective dose ?
Occupational and patient exposure as well as image quality for full spine examinations with the EOS imaging system; Damet et al.

## Slide 19: Exercise
What is the effective dose ?
Occupational and patient exposure as well as image quality for full spine examinations with the EOS imaging system; Damet et al.
thorax
311 mGy.cm2
0.313 mSv
! Units
Radiography
Conversion factor eDAP [mSv/(Gy∙cm2)]

## Slide 20: Solution and comparison
What is the effective dose ?
1.150 x 0.16 + 0.980 x 0.13 = 0.311 mSv

## Slide 21: Automatic exposure control (AEC)
Optimal choice of technical parameters
  kV, mA
  to optimize patient dose and image quality
Radiation detector behind (or in front of) the image detector
Exposure is terminated when the required dose has been integrated

## Slide 22: Summary
Focus (x-ray source)
Tube housing
Filter
X-ray beam
Patient
Antiscattering grid
Détecteur
(AEC)

## Slide 23: Automatic exposure control
In Automatic Exposure Control (AEC) mode, the current and the exposure time are chosen automatically by the system. X-ray installations are equipped with a detector dose measurement system (cell) to stop irradiation as soon as a sufficient dose has reached the detector.
The choice of the position of the cell is therefore critical for the dose delivered to the patient.

## Slide 24: Automatic exposure control
Most installations are equipped with 3 cells: one central and two lateral.
  If a cell covers an area of very absorbent tissue (bone tissue for example), it will receive little radiation and will prolong the exposure time (high dose). The situation is reversed if the cell is behind a more radio-transparent tissue.

## Slide 25: Contents of the lecture today
Introduction (15 min)
  History of x-ray imaging
  Principles of x-ray imaging
Radiography (1h 15 min)
  Chain of image formation
  X-ray production
  X-ray beam
  Interaction with patient
  Detection systems
  Dosimetry
Image Quality (45min )
  Physical measures of image quality
  Acquisition parameters
  Influence of acquisition parameters on patient dose and image quality

## Slide 26: Image quality?
Parameters of influence?
  contrast
  resolution
  noise
  task
  observer
  viewing conditions
Goal of image quality?
  reproduction of reality
  produce useful info
  easy to read
  comparable to the past
  globality (not just for one image)

## Slide 27: Fryback’s taxonomy: Evidence-Based Radiology
  Level 5: Patient outcome
  Level 4: Therapeutic efficacy
  Level 3: Diagnostic thinking
  Level 2: Diagnostic accuracy
  Level 1: Technical efficacy (image noise, CNR, spatial resolution, etc.)
Societal efficacy
Fryback DG et al. Med Decis Making 1991; 11:88-94
Level 6:
Hierachical model for assessing the efficacy in diagnostic imaging
Notes: Diagnostic efficacy
Diagnostic impact
Therapeutic impact
Contribute to patient care

## Slide 28: Contents of the lecture today
Introduction
  History of x-ray imaging
  Principles of x-ray imaging
Radiography
  Chain of image formation
  X-ray production
  X-ray beam
  Interaction with patient
  Detection systems
  Dosimetry
Image Quality
  Physical measures of image quality
  Acquisition parameters
  Influence of acquisition parameters on patient dose and image quality

## Slide 29: Image quality – Level 1
Image quality : at what level ?
The X-ray beam modulated by the patient
  Differential attenuation by the tissues
Sources of image quality reduction
  Scatter
    Compton*
    Thompson-Rayleigh*
  Blurring
    Focal spot
    Movement*
    Image detector
  Noise
    Quantum noise
    Detector noise
*) patient dependent
Strategy
Detector qualification
  Definition of an ideal detector
Influence of the imaging chain
  Contrast reduction
    Scatter
  Spatial resolution reduction
    Scatter
    Geometric blur
    Motion blur

## Slide 30: Image quality – Level 1
Image quality : at what level ?
The X-ray beam modulated by the patient
  Differential attenuation by the tissues
Sources of image quality reduction
  Scatter
    Compton*
    Thompson-Rayleigh*
  Blurring
    Focal spot
    Movement*
    Image detector
  Noise
    Quantum noise
    Detector noise
*) patient dependent
Physical measures of image quality
Contrast
Noise
Spatial Resolution

## Slide 31: Contrast
Prime importance for object detection
Difference in average signal between an object and its neighborhood
Normalized contrast by mean value of background signal
  Depend no longer on detector gain and dynamic range
  Comparison possible between images
  Expressed as percentage
Distinguish between radiant contrast, image contrast and display contrast

## Slide 32: Radiant contrast
Difference in the amount of radiation between different areas of the image
Created by differences in radiation attenuation in diagnostic radiology
Determined by:
  Difference of linear attenuation coefficients
  Thickness of object
  Energy of radiation (determined by voltage and filtration)
Information of internal structure of object
Contrast

## Slide 33: Exercise: calculate the contrast of the following situation
D
F1
F2
F1 = ?
F2 = ?
d
mwater
mal
Normalized contrast = ?
F0

## Slide 34: Exercise: calculate the contrast of the following situation (Solution, normalized)
D
F1
F2
d
mwater
mal
F0
C=
F1-F2
F1
C = 1 - e-(mal -mwater)d
F2
F1
= 1 -
C = 1 - e-Dmd

## Slide 35: Exercise: Use 70 kV or 140 kV for better contrast ?
Remember contrast depends on :
  Attenuation of tissues (µ)
  Scatter (depending on size of patient, energy)
To improve contrast : work at 70 kV
More photons absorbed by the patient
Less scatter radiation
more contrast
But: more dose to patient !

## Slide 36: Effect of scattered radiation on contrast
Scattered radiation does not contribute to the contrast
  Anti-scatter grid !
Only primary radiation creates contrast
  Grid improves contrast by a factor Tp/Tt (primary Transmission/total Transmission)
Scatter decreases the radiant contrast by a factor of 1 + S/P (S=scatter, P=primary)
Contrast

## Slide 37: Image contrast
Detector: transforms radiant contrast into image contrast formed by differences in pixel average values
Transformation is linear
Absolute image contrast = absolute radiant contrast x detector gain
Detector gain depends on detectors radiation response
Image contrast modified by image processing
Contrast

## Slide 38: Noise
Interaction photons-detector is random phenomenon
  Random fluctuations from pixel to pixel = noise
Planar x-ray imaging: Poisson distribution
Dispersion of pixel value around average value reflects noise level
  Noise level quantified by standard deviation of the distribution of pixel values
Input
Output
Imaging
System

## Slide 39: Additional noise sources
Quantum noise:
  Random detection of x-rays
  Main source of noise
  Increases with Nphotons1/2
Electronic noise:
  Produced by the detection chain (detector, amplification stages, electronic circuits, digitizer, etc.)
  Significant contribution to low dose noise
Structured noise:
  Arises from granularity and inhomogeneities of the detector
  Static noise
  Not dependant on dose at detector
Noise

## Slide 40: Noise characterization
In the image domain
  Standard deviation
In the frequency domain
  Noise power spectrum
Noise

## Slide 41: Noise characterization: Standard deviation

First order statistics
  Gaussian distribution around mean value
  Distribution follows a Poisson distribution
  Tends towards a normal distribution symmetrical to average pixel value
1600
1200
800
400
0
# of pixels
1100
1050
1000
950
900
Gray level
Noise

## Slide 42: Noise characterization: Noise Power Spectrum
Noise transfer from image domain in frequency domain
NPS = Fourier transform of the noise autocorrelation
Gives intensity of noise as a function of spatial frequency
Frequency decomposition of the variance
Normalised noise power spectrum
Noise
Notes: as a more thorough noise descriptor than pixel standard deviation, describes both the magnitude and texture as it describes the spatial frequency dependence of the noise

## Slide 43: Noise characteristics: white noise
NPS
1040
1000
960
250
200
150
100
50
0
1600
1200
800
400
0
# of pixels
1100
1050
1000
950
900
Gray level
Noise

## Slide 44: Noise characteristics: low frequency noise
NPS
1004
1000
996
992
250
200
150
100
50
0
3000
2000
1000
0
# of pixels
1020
1010
1000
990
980
Gray level
Noise

## Slide 45: Noise characteristics: high frequency noise
NPS
20
0
-20
200
150
100
50
0
800
600
400
200
0
# of pixels
40
20
0
-20
-40
Gray level
Noise

## Slide 46: Noise characteristics and detection
Same low contrast target in two noise structures
For a given standard deviation noise structure modifies the low contrast target detection
Noise

## Slide 47: Signal detection: Signal-to-noise ratio
SNR: Signal-to-noise ratio
Ratio between signal level (average pixel value) and noise level (standard deviation)
High SNR: good image quality
Noise

## Slide 48: Noise characteristics: Quantum noise
Quantum noise:
2  number of absorbed photons (f)
Signal  number of absorbed photons (f)
Signal-to-noise ratio : S/s  f1/2
1 photon
/ pixel
10 photons
/ pixel
100 photons
/ pixel
1000 photons
/ pixel
Noise

## Slide 49: Spatial resolution
Input
Output
Imaging
System
Definition: ability to distinguish (resolve) small objects or two nearby structures in the image
Signal scattering phenomena
  spread out the signal received
  blur
Resolving power or sharpness
Notes: Resolution is an imaging system's ability to distinguish object detail

## Slide 50: Spatial resolution
Measured signal
Image
Dispersion function of a point
Measured signal
Definition: ability to distinguish (resolve) small objects or two nearby structures in the image
Signal scattering phenomena
  spread out the signal received
  blur
Resolving power or sharpness

## Slide 51: Spatial resolution characterization
In the image domain
  Line spread function (LSF)
  Edge spread function (ESF)
  Point spread function (PSF)
In the frequency domain
  Modulation transfert function
    Spatial frequency response of an imaging system or a component
    Capacity of the detector to transfer the modulation of the input signal as a given spatial frequency to its output
    Contrast as a given frequency relative to low frequencies
    Use of a test pattern
      Deconvolution required (square wave)
    Use of PSF, LSF, ESF
Spatial resolution
Notes: Both resolution and contrast are essential for an image to appear sharp.
In simple terms, resolution is the ability to transfer details, while contrast is the ability to distinguish between light and dark areas. Even if resolution is high, a low contrast will inhibit your ability to clearly see sample details.
MTF measures the ability of the system to transfer the contrast of a sample to an image using spatial frequency (resolution) –> is often expressed in terms of line-pairs per millimeter = frequency
As line spacing decreases (i.e. the frequency increases) , it becomes increasingly difficult to efficiently transfer this decrease in contrast;
At higher spatial frequency MTF falls towards 0 , this corresponds to the poor visibility of the small structures. Conversely, at lower spatial frequency MTF is closer to 1 and it represents the ability to clearly visualize large structures.

## Slide 52: In practice: measure spatial resolution
Image
Profile
X-ray
L(mm) = FWHM
Object : different spatial frequencies
Image of the object
MTF ability of system to transmit a frequency band
Image domain:
Frequency domain:
Spatial resolution

## Slide 53: Response to an impulse function
Response to an punctual impulse : PSF
  Image of a pin hole
    Requires high dose
    Produce a weak signal
Response to a linear impulse : LSF
  Image of a slit
Response to an edge input : ESF
  Image of an edge
Spatial resolution

## Slide 54: MTF
Resolution of imaging system often described in frequency space (MTF)
Resolution pattern: metal bars at different spatial frequencies
The higher the frequencies (the closer the bars) the more the attenuation of the signal
MTF: amplitude of periodic signal of bars as a function of the pairs of lines
object
image
Intensity profiles
Spatial resolution

## Slide 55: Response to an impulse function
Spatial resolution

## Slide 56: In practice : Impulse function (Example)
Edge response (ESF)
Linear impulse response (LSF)
MTF
Spatial resolution

## Slide 57: In practice: Example of application
Spatial resolution
Notes: it is the contrast at a given spatial frequency relative to low frequencies
By definition, the low frequency MTF limit is always 1 (100%)

## Slide 58: Limiting factors for spatial resolution
Geometric factors  geometric blur (penumbra)
  The optical size of the focal spot of the x-ray tube matters
  Size of penumbra depends on magnification factor
  Avoid geometric blurring: focus size in image should be smaller than pixel size
Spatial resolution

## Slide 59: Limiting factors for spatial resolution
Geometric factors  geometric blur (penumbra)
Pixel size = (size of field of view)/(size of image matrix)
  The smaller the pixel the better the spatial resolution
  Maximum frequency visible in an image (cutoff frequency): fcut = 1/(pixel size)
Spatial resolution

## Slide 60: Limiting factors for spatial resolution
Geometric factors  geometric blur (penumbra)
Pixel size = (size of field of view)/(size of image matrix)
Powder detector
Needle detector
Signal scatter in detector
  The thicker the detector layer the greater the scatter
  Direct conversion DR systems do not produce light  advantageous
  Needle-structure: limits scatter (channel light)
Spatial resolution

## Slide 61: Limiting factors for spatial resolution
Geometric factors  geometric blur (penumbra)
Pixel size = (size of field of view)/(size of image matrix)
Signal scatter in detector
Motion blur (temporal resolution)
  Movement during exposure time creates blur
  Size of blur proportional to movement speed and exposure time
Extremities | 0.5 – 1
Head | 1 - 2
Abdomen sup | 20 – 40
Abdomen inf | 15 – 30
Lungs (adults) | 75 – 100
Lungs (children) | 150 – 200
Cardiac muscle (adults) | 75 – 100
Cardiac muscle (children) | 150 – 200
Heart | 400 - 500
[mm/s]
Chest X-ray without breath hold (5 ms exposure time):
BM = 75 x 0.005 = 0.375 mm
Spatial resolution

## Slide 62: Limiting factors for spatial resolution
Geometric factors  geometric blur (penumbra)
Pixel size = (size of field of view)/(size of image matrix)
Signal scatter in detector
Motion blur (temporal resolution)
  Movement during exposure time creates blur
  Size of blur proportional to movement speed and exposure time
Extremities | 0.5 – 1
Head | 1 - 2
Abdomen sup | 20 – 40
Abdomen inf | 15 – 30
Lungs (adults) | 75 – 100
Lungs (children) | 150 – 200
Cardiac muscle (adults) | 75 – 100
Cardiac muscle (children) | 150 – 200
Heart | 400 - 500
[mm/s]
Chest X-ray without breath hold (5 ms exposure time):
BM = 75 x 0.005 = 0.375 mm
Spatial resolution

## Slide 63: Contents of the lecture today
Introduction
  History of x-ray imaging
  Principles of x-ray imaging
Radiography
  Chain of image formation
  X-ray production
  X-ray beam
  Interaction with patient
  Detection systems
  Dosimetry
Image Quality
  Physical measures of image quality
  Acquisition parameters
  Influence of acquisition parameters on patient dose and image quality

## Slide 64: Chain of image formation

## Slide 65: Chain of image formation
X-ray tube
Voltage [kV]
Current [mA]
Time [s]
Dose
Noise
Attention: kV ↑ contrast ↓

## Slide 66: Chain of image formation
X-ray tube
Filtration
adjust
Filtration
Dose

## Slide 67: Chain of image formation
X-ray tube + Filtration
Field of View
FOV
Dose
Contrast
Noise
Artefacts

## Slide 68: Chain of image formation
X-ray tube + Filtration
Field of View
Detector
Detector performance
Dose
Image
Quality

## Slide 69: Chain of image formation
X-ray tube + Filtration
Field of View
Detector
Anti-scatter grid
Use of anti-scatter grid
Dose
Image
Quality

## Slide 70: Chain of image formation
X-ray tube + Filtration
Field of View
Detector + anti-scatter grid
Pixel size
Pixel Size
Dose
Spatial
resolution
Noise

## Slide 71: Mammography

## Slide 72
X-Ray tube
Compression plates

## Slide 73: High image quality requirements
The radiological mammography examination requires high image quality (good contrast and spatial resolution) and strict control of the dose:
The main objects of interest to highlight in the breast are:
  microcalcifications (small size ~ 0.1 mm)
  tumor masses of low contrast (size 5 to 10 mm).

## Slide 74: High image quality requirements
The nature of the breast tissue is highly variable, with relatively dense parenchymal tissue or with fatty tissue.
The breast is a radiosensitive organ: WT = 0.12
  it corresponds to 12% of the risk of death from cancer in the event of homogeneous whole-body irradiation

## Slide 75: Difference between mammography and radiography
Devices used in mammography differ from conventional x-ray devices:
The voltage used in mammography is low to produce good contrast between tissues
  between 25 and 40 kV (between 50 and 120 kV in radiography)
  possible because the breast is thin (2 to 9 cm).

## Slide 76: Radiation dose versus energy
0.1
1.0
10
100
Skin dose (mGy)

## Slide 77: Optimum energy vs thickness
To get reasonable dose level SNR will have to be reduced as a function of breast thickness
Image quality has to be checked at various thicknesses

## Slide 78: Difference between mammography and radiography
The half-value layer (HVL) is small,
  order of 0.4 mm aluminum (2.5 mm in radiology).
The radiation spectrum is more mono-energetic
  predominance of x-ray characteristics of molybdenum (Mo) or rhodium (Rh) with a filter with Mo (use K-edge filter)

## Slide 79: Difference between mammography and radiography
The focal spot of a mammography tube is small compared to a conventional x-ray tube
  about 0.3 mm for the large focus and 0.1 mm for the small focus.
Tube - detector distance is shorter,
  around 65 cm in mammography and 100 cm in radiography

## Slide 80: Exercise
What is the magnification for these two configurations ?
Data
SID = 60 cm
SOD = 54 cm then 50 cm
OID = 6 cm then 10 cm
F = 0.3 mm

## Slide 81: Solution
What is the magnification for these two configurations ?
Data
SID = 60 cm
SOD = 54 cm then 50 cm
OID = 6 cm then 10 cm
F = 0.3 mm

## Slide 82: Magnification
Magnification: improve the analysis of microcalcifications
  Important that the focus is small if we want keep good resolution.
The dose at the breast is however higher when using magnification and increases according to the square of the magnification factor.

## Slide 83: Difference between mammography and radiography
Mammography tubes are less powerful with a lower heat capacity than conventional x-ray tubes.
  generally 100 mA on the large focus and 25 mA on the small focus.
In mammography, exposure time is much longer than in conventional radiography
  exposure times of 0.5 to 2 s are frequently used.
Diagnostic screens used in mammography must have a higher luminance range
  at 250 cd/m2 while those for conventional radiography are satisfactory with a range of luminance of 150 cd/m2

## Slide 84
In practice, the choice of voltage is made by the automatic exposure system (AEC) according to breast thickness and density.
The x-ray of a thick or dense breast loses in contrast, but the choice of voltage must favor the best contrast-dose compromise.
On the machine, the choice of voltage is made according to the height of the compression plate and / or a measurement of breast attenuation during a preliminary shot of a small amount of X-rays before the realization of the exam.

## Slide 85: Compression
Adequate compression of the breast is one of the most important conditions for quality mammography and for a good visualization of possible pathologies.
Improved Subject Contrast
  Better penetration allows use of lower kVp
  Less Scatter (thinner patient)
Better Resolution
  Less geometric blur: objects pressed closer to film
  Less motion: patient immobilization
  Better detection by spreading out structures
Better Conspicuity: less tissue/object overlap
Lower Radiation Dose: less mAs needed
Bad compression
Good compression

## Slide 86: Motion Artefact
With movement
Without movement
Notes: https://www.youtube.com/watch?v=5bsqhY_PtkA

## Slide 87: Control of scattering
Breast compression
  Spreads out breast anatomy
    Reduces superposition
  Reduces scatter
  Reduces motion blur
  Reduces absorbed dose
Suspicious  normal with
Spot compression
Scatter to primary ratio : 0.8-1.0
Scatter to primary ratio : 0.4-0.5
Not optimal
Optimal compression

## Slide 88: To improve low contrast lesion visibility
Increase of the Contrast to Noise ratio (CNR)
  To increase contrast:
    Use lower beam energy
      kVp
      Anode/filter
       Radiation dose will increase (less penetrating)
  To reduce quantum noise
    increase the dose to the image detector
  Often a slight increase of beam energy allows for a slight contrast reduction but a noticeable increase of the dose to the detector  CNR improvement

## Slide 89: Anti-scatter grid
In mammography, the contribution of scattered radiation is less important than conventional radiology
  the thickness of the compressed breast is relatively small (2 to 9 cm) and the field area is limited (generally 18 x 24 cm2).
  The transmission of primary radiation is approximately 70% and that of the scattered radiation of the order of 20%.

## Slide 90: Special features of breast tissue
Breast tissue is composed with glandular tissue and fatty tissue.
The proportion of glandular tissue determines the rate of glandularity
  decreases with age and breast thickness.
  The coefficient attenuation of adipose tissue is lower than that of the denser glandular tissue.

## Slide 91: Exercise
What is the half value layer at 20 keV and 50 % / 50 % tissue?
Energy [keV]
Attenuation coefficient [cm-1]
Adipose tissue
50%/50% tissue
Glandular tissue

## Slide 92: Exercise
What is the half value layer at 20 keV and 50 % / 50 % tissue?
HVL
Energy [keV]
Adipose tissue
50%/50% tissue
Glandular tissue
Attenuation coefficient [cm-1]

## Slide 93: Subjective image quality assessment
The number of parameters involved in a mammogram is important.
Optimization of settings is therefore difficult.
Phantom objects can be used to simulate a standard breast (thickness and composition) which contains objects simulating microcalcifications, masses and filaments of different sizes.
  it possible to establish a detection score (detectability index)

## Slide 94: Subjective image quality assessment
European institutions establish minimum contrast visibility thresholds according to the size of the objects.
Filaments
Calcifications
Masses

## Slide 95: Dosimetry in mammography

## Slide 96: Entrance Surface Dose (ESD)
Dose measured at the surface of the breast at the reference point 6 cm from the chest wall
For a given dose to the detector, the dose increases exponentially with the thickness of the breast, hence the essential role of judicious choice of voltage and compression
    C : radiological constant
    U : voltage [kV]
    Q : charge
    d : distance
Dose (mGy)
Compressed breast thickness (mm)

## Slide 97: Mean glandular dose (MGD)
The mean glandular dose is representative of the radiological risk
The dose distribution is not uniform
  in the breast; 50% reduction per cm
Limit values of MGD must be respected by systems in screening for different breast thicknesses
    g : takes into account the transmission of X-rays for a breast of average composition 50% glandular - 50% fatty
    c : corrects the g factor for other glandularity levels
    s : takes into account the quality of the spectrum related to the anode / filter pair