## Slide 1
FLUOROSCOPY
FLUOROSCOPY
Dre Marie Nowak
Institut of Radiation physics (IRA) – UNIL – CHUV
marie.nowak@chuv.ch
Master of Science EPF-ETH degree in Nuclear Engineering
Medical radiation physics

## Slide 2
INTRODUCTION
INTRODUCTION
In the medical field, it is common to need to visualize the inside of the human body.

## Slide 3
INTRODUCTION
INTRODUCTION
In the medical field, it is common to need to visualize the inside of the human body. There are two options for this:
First: Surgery

## Slide 4
INTRODUCTION
INTRODUCTION
Second: Imaging

## Slide 5
INTRODUCTION
INTRODUCTION
There are many imaging techniques (for diagnostic or therapeutic use). Radiography, CT, MRI, fluoroscopy… We will focus here on fluoroscopy.
This technique is increasingly used as an alternative to surgery. Indeed it allows the radiologist to move a catheter in the human body and to visualize it as the intervention progresses.
Fluoroscopy is less invasive than surgery (only one entry point into an artery). Fewer complications are observed, and hospitalization time is reduced. Procedures are becoming more and more complex, with increasingly long irradiation times, which can lead to high irradiation of patients and staff.

## Slide 6
INTRODUCTION
INTRODUCTION
Fluoroscopy is only rarely used for diagnosis unlike CT or radiography. It is mainly used for the treatment of patients.
Diagnosis: Biopsy
Treatment: - Stroke
- Stent placement
- Cancer treatment (chemo administration)

## Slide 7
Radiography et fluoroscopy
Similarities
Projection of the anatomy
Similar X-ray kV range but beam energy might vary
Image detector: different for analog/similar for digital
INTRODUCTION
Differences
Recording of moving structures (dynamic)
Order of magnitude of patient/staff exposure
Fluoroscopy requires puncture (vessel entry)
Diagnostic/therapeutic

## Slide 8
Reminder
Radiological projection technique allowing visualization of the anatomy in real time using X-rays
As for the x-ray, the contrast depends on:
Atomic number of tissues (photoelectric)
Radiation energy
The thickness of the anatomical structures
Tissue density (Compton)
Irradiated volume
Compton => scattered radiation production
Image contrast
Radiation protection of workers
INTRODUCTION
aneurism

## Slide 9
FLUOROSCOPY
FLUOROSCOPY SYSTEM

## Slide 10
Fluoroscopy
FLUOROSCOPY
What is it? How does it work?

## Slide 11
Fluoroscopy
FLUOROSCOPY

## Slide 12
Fluoroscopy
FLUOROSCOPY
Nowadays: two detector technologies to give a digital image:
Flat panel detector
Luminance amplifier
Analog technology: Signal digitization at its output or for new machines: digital technology
Direct detector with indirect conversion (see digital radiography)

## Slide 13
Luminance amplifier
LUMINANCE AMPLIFIER
Image Intensifier
Conversion of X-ray into visible light
Conversion of light photons into electron
Electron acceleration and focalisation
Conversion of the accelerated electrons into visible light
Gain of « minification »
Large image at the entrance and production of a small image in the output
Gain of « flow »
Electron acceleration increases the amount of light produced by the output scintillator

## Slide 14
Luminance amplifier
LUMINANCE AMPLIFIER
Entrance of the image intensifier
Photon light hit the photocathode
photoelectric effect
= electrons
Incident photon
Photoelectric or Compton effect = Light photon

## Slide 15
Luminance amplifier
LUMINANCE AMPLIFIER
Image intensifier - details
Entry CsI (Cesium Iodide)
  Good efficiency at kV used (Kedge 35 - 40 keV)
  Emission of 420 nm photons
    Spatial resolution - Efficiency compromise
    CsI under the form of crystal needles (5 µm : light guide)
    Thickness of 300 to 450 µm
Photocathode (CsSb alloy)
    10 - 20 % of visible light into electrons
Acceleration through 25 to 35 kV
Silver-activated Zinc-Cadmium sulphide (Zn Cd S: Ag) : P20
  about 400 photo-electrons for 1 photon of 60 keV
  Emission of 520-540 nm photons
    Thickness 4 to 8 µm; size of the crystals 1 to 2 µm
    1 electron of 30 keV produces 1000 light photons
Optical coupling to a camera (with apertures that can be varied)

## Slide 16
Flat Panel detector
FLAT PANEL DETECTOR
Image intensifier - details
The needle architecture behaves like an optical fiber and directs the light to the photodiodes. This improves the spatial resolution compared to an amorphous structure.
Pixel size: 140 – 200 mm See Dr. Vitzthum's course.

## Slide 17
Luminance amplifier – Flat panel detector
AMPLIFIER VS FPD
 | Luminance amplifier | Flat panel detector
 | Analog or Digital | Digital
FOV | Small to big | Small to very big (30 cm x 40 cm)
Image per sec | 1 to 25 ips - continuous | 0.25 to 60 ips
Price | Cheaper | More expensive
Volume | Can be moved to an other room Big detector | Fixe Compact
Application | Operating room – vascular or spinal surgery Pneumology | Interventional radiology Cardiology
NB: FOV = field of view
IMAGE

## Slide 18
FLUOROSCOPY SYSTEM
Fluoroscopy system: the basics
Two operating modes:
Fluoroscopy
Fluorography

## Slide 19
FLUOROSCOPY SYSTEM
Fluoroscopy system: the basics
Two operating modes:
Fluoroscopy
Fluorography
Low dose
High filtration (copper)
Low image quality
Noisy image
Used during displacement of the catheter
High dose
Very good image quality
Used to saved the images
Fluorography mode should be used with care, always prefer Fluoroscopy mode
kV and mA are adapted increase of kV: low dose mode
mA increase: increase dose
Brand dependant

## Slide 20
FLUOROSCOPY SYSTEM
Fluoroscopy system: the basics
Fluoroscopy
Fluorography
5 to 50 mGy/min (to 15 ips)
NB: Data for 20 cm PMMA
DSA:
0.5 to 5 mGy/image
1 to 5 ips – run of 30 images – few seconds
= 15 to 150 mGy per sequence
Cine : 50 to 150 mGy/min (to 15 ips)

## Slide 21
FLUOROSCOPY SYSTEM
Fluoroscopy system: the basics
Effect of copper filter
Without Cu filter
Without Cu filter
X-ray spectrum at the entrance and exit of the patient – (without Cu filter)

## Slide 22
FLUOROSCOPY SYSTEM
Fluoroscopy system: the basics
Automatic exposure control (AEC)
Continuous monitoring of light output
Measurement of the average signal received by the detector
Action on the generator
 kV mA/ms per image

## Slide 23
FLUOROSCOPY SYSTEM
Fluoroscopy system: Fluorography mode
Digital substraction angiography: DSA

## Slide 24
FLUOROSCOPY SYSTEM
Fluoroscopy system: Fluorography mode
Digital substraction angiography: DSA
The images are subtracted but the noise is added so the dose must be significantly increased in this operating mode

## Slide 25
FLUOROSCOPY SYSTEM
Fluoroscopy system: Fluorography mode
Road mapping

## Slide 26
FLUOROSCOPY SYSTEM
Fluoroscopy system: Fluorography mode
Road mapping
Very useful for positioning catheters or wires in complex geometries or small vessels.
They start with the DSA, the most opacified image is kept and will serve as a basis for the following images which will be superimposed on this mask. We see the guide or the catheter evolving on a static opacified image

## Slide 27
FLUOROSCOPY SYSTEM
Fluoroscopy system: Tips to save dose
Last Image Hold: LIH
The last frame before stopping x-ray is continuously displayed

## Slide 28
FLUOROSCOPY SYSTEM
Fluoroscopy system: Tips to save dose
Recursive filters
The dose rate is reduced but the temporal resolution of the acquisition is also reduced (increase in the correlation between the images presented). Motion blur induced.

## Slide 29
FLUOROSCOPY SYSTEM
Fluoroscopy system: Tips to save dose
Edge enhancement
Big image treatment leds to decrease dose, because better image quality

## Slide 30
Image quality – Differences between both systems
IMAGE QUALITY
Dynamic range
Dazzled or dark areas
Grey levels
Detector air Kerma
Artery
Spine
Lung
Low dynamic range
Image with luminance amplifier
Image with FPD
Artery
Spine
Lung
Large dynamic range
Detector air Kerma
Grey levels
Increased discrimination of the vessels compared to the anatomical background

## Slide 31
Image quality – Differences between both systems
Dynamic range
Use of semi-transparent filters to equalize the fluence at the input of the luminance amplifier
IMAGE QUALITY

## Slide 32
Image quality – Differences between both systems
Luminance amplifier limitation
Distorsion / vignetting
IMAGE QUALITY

## Slide 33
Image quality – Differences between both systems
Luminance amplifier limitation
Distorsion / vignetting
« Pin cushion » artefact (warping)
« Vignetting »
  Distortion due to the geometry of the luminance amplifier
  Effect that increases with the size of the field
IMAGE QUALITY

## Slide 34
Image quality – Differences between both systems
Luminance amplifier limitation
Distorsion / vignetting
IMAGE QUALITY

## Slide 35
Image quality – Differences between both systems
Luminance amplifier: electronic magnification
NB: FOV = field of view
Electronic magnification:
For the same signal / noise ratio, the same number of photons is required at the output of the amplifier
The dose increases as the area of ​​the field size decreases (field size divided by 2  dose rate increase by 4)
Reduction of the signal-to-noise ratio at magnification
In electronic magnification, the gain in minification is reduced  the dose rate must be increased
IMAGE QUALITY

## Slide 36
Image quality – Differences between both systems
Flat panel detector: Pixel binning
NB: FOV = field of view
To save dose,
Reduce the size of the image,
And reduce the noise level in the image
Pixel binning
However, it reduces the spatial resolution
IMAGE QUALITY

## Slide 37
Image quality – Differences between both systems
Flat panel detector: Pixel binning
IMAGE QUALITY

## Slide 38
Image quality – Differences between both systems
Constrate resolution
NB: FOV = field of view
Numbers of bits inluence grey scale so contrate resolution
IMAGE QUALITY

## Slide 39
Image quality – Differences between both systems
Magnification and dose rate
NB: FOV = field of view
IMAGE QUALITY

## Slide 40
AMPLIFIER VS FPD
Summary
Luminance amplifier – Flat panel detector
 | Luminance amplifier | Flat panel detector
Cost | Cheaper | More expensive
Space | Voluminous | Compact
Image per sec | 1 to 25 ips - continuous | 0.25 to 60 ips
Image quality | Image distorsion Vignetting Magnification - Dose | 
Dose reduction | Egal to FPD | Egal to Amplifier
Amplifier
FPD

## Slide 41
FLUOROSCOPY
RADIATION PROTECTION

## Slide 42
Radiation protection ? Why ?
Because radiation have body impacts:
Reminder
  Stochastic effects
    Ex: Cancer
    No threshold
    Concerns: Workers (and patients)
  Deterministic effects
    Ex: Skin burn
    Threshold for skin burn: 2 Gy
    Concerns: Patients
We want to avoid unnecessary radiation
RADIATION PROTECTION

## Slide 43
Radiation protection ? Why ?
RADIATION PROTECTION

## Slide 44
Radiation protection
RADIATION PROTECTION
Dose Limits
Exposed worker
(Adult)
Public & pregnant woman
Patient
No dose limit is set for medical exposure
Whole body | 20 mSv/year
Eye lens | 20 mSv/year
Extremities | 500 mSv/year
1 mSv/year
Need of radiation protection

## Slide 45
Radiation protection
RADIATION PROTECTION
Choose your priorities: act first
Completion of the procedure is of course a priority over radiation protection, it is "rare" for a doctor to deem the radiological risk sufficiently great not to intervene under X-ray
The radiological risk is often faced with clinical certainty
        Success of the interventional act
        Radiation protection

## Slide 46
Radiation protection
RADIATION PROTECTION
How can we know if an exam was too irradiating ?

## Slide 47
Radiation protection
RADIATION PROTECTION
Diagnostic reference levels (DRL)
For a given examination, corresponds to the 75th percentile of the distribution of a dosimetric indicator (for exemple dose).
Collected, updated and edited by the FOPH.
It is not a limit, but if in a center all the exams are above this limit, a control of the device is needed. On the opposite, if dose are very far (under) the DRL, the practician can miss something. It is important to check the device to be sur that all clinical information is present on the images with such a small dose

## Slide 48
Radiation protection
RADIATION PROTECTION
Diagnostic reference levels (DRL)

## Slide 49
Radiation protection
RADIATION PROTECTION
Diagnostic reference levels (DRL)
4 fundamental quantities

## Slide 50
Radiation protection
RADIATION PROTECTION
Fluoroscopy: 4 fundamental quantities
Kerma area product (KAP or PDS)
Fluoroscopie time
Air Kerma cumulated
= skin dose
Number of images
At the end of each exam, we obtain a dose report:

## Slide 51
Fundamental quantities – Fluoroscopie Time
FUNDAMENTAL QUANTITIES
Fluoroscopie time:
107 min
< 2 Gy
Alarm at 90 min
Do not necessarily stop the exam at 90 min
Look at the Kerma, are we far from 2 Gy, and 5 Gy?
Is the exam still long? Can another more experienced operator continue ?

## Slide 52
Fundamental quantities – Fluoroscopie Time
Fluoroscopie time:
Advantage:
Representative of the complexity of an intervention
Disadvantages:
Not very representative of the effective duration of radiation in pulsed fluoroscopy.
Example :
10 minutes of fluoroscopy → 600 seconds
Pulsed scopy at 15 frames per second
One image: ~ 10ms, or 0.01s
“Real” exposure time: 15 fps x 0.01 s / im. X 600s
90 seconds of beam in reality, or 15% of the time
Do not take into account the size of the patient, the acquisition mode (scopie low dose, high dose…)
FUNDAMENTAL QUANTITIES

## Slide 53
Fundamental quantities – Number of images
FUNDAMENTAL QUANTITIES
Number of images
Indication on the complexity of the examination and the patient exposure
A radiographic image generally requires a much higher dose rate (reduced noise in the image)
Only diagnostic images (graphics, DSA) are saved
Scopy, usually no backup
Some installations allow manual saving, useful for reviewing the sequence (without emitting x-rays)

## Slide 54
Fundamental quantities - Kerma
Kerma [mGy]
The Patient Entrance Dose (ED), is the dose in air at a point on the patient's surface. Problem : It is determined by the dose at the Interventional reference point (IRP). The reference point is fixed at 15 cm from isocenter.
FUNDAMENTAL QUANTITIES

## Slide 55
Fundamental quantities - Kerma
Kerma [mGy]
Tends to overestimate dose in thin patient  might be better
Tends to underestimate dose in large patient  not good!
FUNDAMENTAL QUANTITIES

## Slide 56
Fundamental quantities – KAP
Kerma area product (KAP) or Dose area product (DAP)
FUNDAMENTAL QUANTITIES
Distance
Dose
Surface
KAP
  Be careful on the units used
    Can be measured or calculated on the unit
    Can take into account the table attenuation
Independent of distance to the source
Indicative of the risk of stochastic effects for the patient
Standard unit: Gy.cm2

## Slide 57
Fundamental quantities
Peak skin dose
Obtained from Kerma and tube angulation. Ideal solution but lot of errors possible actually
FUNDAMENTAL QUANTITIES

## Slide 58
Fundamental quantities
Peak skin dose
Obtained from Kerma and tube angulation. Ideal solution but lot of errors possible actually
FUNDAMENTAL QUANTITIES

## Slide 59
Fundamental quantities
Peak skin dose
Obtained from Kerma and tube angulation. Ideal solution but lot of errors possible actually
FUNDAMENTAL QUANTITIES

## Slide 60
Fundamental quantities
Peak skin dose
Obtained from Kerma and tube angulation. Ideal solution but lot of errors possible actually
FUNDAMENTAL QUANTITIES

## Slide 61
Fundamental quantities
Peak skin dose
Obtained from Kerma and tube angulation. Ideal solution but lot of errors possible actually
FUNDAMENTAL QUANTITIES

## Slide 62
Fundamental quantities
Peak skin dose
Obtained from Kerma and tube angulation. Ideal solution but lot of errors possible actually
Which part of the epidermis was the most irradiated?
FUNDAMENTAL QUANTITIES

## Slide 63
Fundamental quantities
Peak skin dose
Obtained from Kerma and tube angulation. Ideal solution but lot of errors possible actually
Which part of the epidermis was the most irradiated?
Peak Skin Dose
FUNDAMENTAL QUANTITIES

## Slide 64
Fundamental quantities
Peak skin dose
Obtained from Kerma and tube angulation. Ideal solution but lot of errors possible actually
Peak Skin Dose
To be accurate, the software need to be correlated to
Patient position on the table
Distance source patient
FOV
Tube angulation
Biplane system
We continue to use the Kerma
Kerma > 2 Gy
FUNDAMENTAL QUANTITIES

## Slide 65
Fundamental quantities
Practical case:
FUNDAMENTAL QUANTITIES

## Slide 66
Fundamental quantities
Skin burn:
Radio-induced burns appear from 2 Gy
  A long procedure,
  A corpulent patient,
  An irradiation in LAT increase the risks
It is necessary :
  To change the angle of incidence;
  Warn the operator when the dose is too high in order to quickly finish the procedure or ask for help from a «senior»
The patient should be warned and monitored for a dose greater than 5 Gy
FUNDAMENTAL QUANTITIES

## Slide 67
Fundamental quantities
Summary:
FUNDAMENTAL QUANTITIES
Short term
  Deterministic effect
  No surprise
Dose indicator
  Kerma at IRP
    Representative of:
      Peak skin dose
Long term
  Stochastic
  To be optimized
Dose indicator
  KAP
    Representative of:
      Effective dose

## Slide 68
FLUOROSCOPY
RADIATION PROTECTION
For the patient

## Slide 69
Radiation protection to the patient
PATIENT RADIATION PROTECTION
How can we decrease the patient exposure ?

## Slide 70
Radiation protection to the patient
Major rules: patient positionning
SEE = ?
Magnification ?
SEE = ?
Magnification ?
Which position do you choose?
N°1
N°2
N°3
They are the same
PATIENT RADIATION PROTECTION

## Slide 71
Patient radiation protection
Tube positionning
PATIENT RADIATION PROTECTION
Image detector
As close as possible to the patient
X ray tube
As far as possible to the patient

## Slide 72
Patient radiation protection
Influence of patient thickness & different modes of fluoroscopy and graphy
PATIENT RADIATION PROTECTION
The dose rate to the patient's skin increases with the thickness of the patient.
The dose rate to the patient's skin increases with the different fluoroscopy modes and with the fluorography
Be careful, in radiography mode, with a patient of 20 cm thick, the 2 Gy on the skin will be reached in 16 min for this example.
Patient thickness [cm]
Patient thickness [cm]
Skin dose rate [mGy.min-1]

## Slide 73
Patient radiation protection
PATIENT RADIATION PROTECTION
1 seconde is represented
X-ray emission fraction during fluoroscopy use
For 1 min of fluoroscopy use
1 min of pulsed fluoroscopy (15 f/s)
~ 9 sec of irradiation
1 min of pulsed fluoroscopy (3 f/s)
~ 2 sec of irradiation
1 min of pulsed fluoroscopy (1 f/s)
~ 500 milli sec of irradiation
Pulsed mode (appropriate choice of protocol)

## Slide 74
Patient radiation protection
PATIENT RADIATION PROTECTION
Pulsed mode (appropriate choice of protocol)
1 – 3 pulses/s
Appropriate for slow processus
Orthopedic protocole
Source : T.A. Zachos, Michigan State University, Kalamazoo, NAVC Conference 2013 Small Anima
Cardiac mouvements
8 – 30 pulses/s
Appropriate for fast processus

## Slide 75
Patient radiation protection
PATIENT RADIATION PROTECTION
Diaphragme
The diaphragm is a key tool: it makes it possible to increase the image quality and lower the doses

## Slide 76
Patient radiation protection
PATIENT RADIATION PROTECTION
Zoom – Tube positionning – geometric zoom
Tube close to the patient, detector far away
Magnified and blurred image
The dose to the skin increases
The detector is close to the patient, the tube far away
Image net
Lower skin dose

## Slide 77
Patient radiation protection
PATIENT RADIATION PROTECTION
Zoom
The geometric zoom consists of bringing the patient closer to the tube
Any installation has an electronic zoom, internal to the machine
The zoom utilization should be limited as the patient is more irradiated.
Between electronic and geometric zoom, it is better to choose electronic zoom. Had to be remove when it's no longer needed.
Skin dose rate [mGy.min-1]
Magnification

## Slide 78
Patient radiation protection
PATIENT RADIATION PROTECTION
Use of patient shielding

## Slide 79
Patient radiation protection
PATIENT RADIATION PROTECTION
Summary
Geometry
Patient far from tube
Beware of oblique incidences
Time
  Reduce fluoroscopy time
  Using the LIH
  Pay attention to the chosen dose mode
  Acquisition frequency
Diaphragme !!!
Moderate use of magnification
Limit the number of recorded images to what is strictly necessary

## Slide 80
FLUOROSCOPY
RADIATION PROTECTION
For workers

## Slide 81
Workers’ radiation protection
WORKERS’ RADIATION PROTECTION
What is irradiating workers?

## Slide 82
Workers’ radiation protection
WORKERS’ RADIATION PROTECTION
The scattered radiation propagates in all direction of space
The scattered radiation is mostly emitted by the patient, (but also by all the materials crossed by the X-ray beam)

## Slide 83
Workers’ radiation protection
WORKERS’ RADIATION PROTECTION
Dose rate around the patient
Challenge: improve the protection of fluoroscopy personnel
Natural dose rate: 0.1 - 0.2 uSv / h
Dose rate near a patient: 0.5 - 4 mSv / h Annual dose limit 20 mSv
Protection by lead apron (transmissions 3 - 4%): but dose to the eye lens
Major problem: predicting particularly dangerous areas

## Slide 84
Workers’ radiation protection
WORKERS’ RADIATION PROTECTION
How can we decrease the worker’s exposure ?

## Slide 85
Workers’ radiation protection
WORKERS’ RADIATION PROTECTION
3 rules of radiation protection
Time
Shielding
Distance
Minimize time near radiation source
Increase distance from radiation source
Dose decrease with distance square
Use shields / screens
Moreover: Formation: Good training in radiation protection

## Slide 86
Workers’ radiation protection
WORKERS’ RADIATION PROTECTION
The patient is the main source
Place the tube under the table

## Slide 87
Workers’ radiation protection
WORKERS’ RADIATION PROTECTION
Effect of incidence
Where to position yourself during lateral irradiation?
Ideally, workers should position themselves on the side of the detector.

## Slide 88
Workers’ radiation protection
WORKERS’ RADIATION PROTECTION
Effect of field size: patient / staff
Passage of a field size of 14 to 28 cm

## Slide 89
Workers’ radiation protection
WORKERS RADIATION PROTECTION
Effect of patient morphology: patient / staff
Passage from 24 to 34 cm thick

## Slide 90
Workers’ radiation protection
WORKERS’ RADIATION PROTECTION
Shielding protection
Collective protection gears:
Table shield/curtain - 98%
Mobile lead screen
Suspended lead screen
-98%
- 98%

## Slide 91
Workers’ radiation protection
WORKERS’ RADIATION PROTECTION
Shielding protection
Collective protection gears:

## Slide 92
Workers’ radiation protection
WORKERS RADIATION PROTECTION
Shielding protection
Collective protection gears:
Gap 5 – 10 cm
Dose rate x 30
Adjusted protections: 0,033 mSv/h
Non adjusted protections: 1 mSv/h
1 mSv/h

## Slide 93
Workers’ radiation protection
WORKERS’ RADIATION PROTECTION
Shielding protection
Personal protection gears:
Lead apron
Collar
Glasses or visor
- 50 to 75%
- 90 to 98%
- 90 %

## Slide 94
Workers’ radiation protection
WORKERS’ RADIATION PROTECTION
Shielding protection
Personal protection gears:
Surgeon (vascular, spinal, traumatology)
Nurse (worker close to surgeon in vascular, spinal, traumat,…)
Surgeon (other spe)
Nurses (other specialities)
Anesthetist
Other workers

## Slide 95
Workers’ radiation protection
WORKERS’ RADIATION PROTECTION
Shielding protection
Personal protection gears:
Ex. Tablier en ordre avec collerette, cf AW07-1
Ex. Tablier déchirure sur couture sangle,
A surveiller
Cf AW07-1
Aprons should be hung up, not thrown down or folded

## Slide 96
Workers’ radiation protection
WORKERS RADIATION PROTECTION
What is wrong in this picture?
Tube
Shielding

## Slide 97
Workers’ radiation protection
WORKERS’ RADIATION PROTECTION
Distance
Patient head (with shieling screen)
2 m away (vs patient head)
Eye lens | - 99 % | - 75 à 90 %
Chest | - 99 % | - 90 à 95 %
Leg | - 99 % | - 90 à 95 %
Dose reduction compare to a worker close to patient head without shieling creen
Reference position

## Slide 98
Workers’ radiation protection
WORKERS’ RADIATION PROTECTION
Distance
One step ÷ 3
Two steps ÷ 10

## Slide 99
Workers’ radiation protection
WORKERS RADIATION PROTECTION
Dosimetry
Whole body dosimeter (mandatory)
Double dosimetry
(depend on function)
From double dosimetry: Eye lens dose
Dosimetry of extremities (depend on function)
Under apron
IRA 2010-09
Above apron
IRA 2010-09

## Slide 100
Workers’ radiation protection
WORKERS’ RADIATION PROTECTION
To control the personnel exposure
Radiation protection of the patient
Distance
As far as possible on fluoroscopy
Use of automatic injectors in acquisition
Protective screen
Pb apron (0.35 mm Pb): 90-98% attenuation
Thyroid protector (0.5 mm Pb): 90%
Mobile screens (1 mm Pb): 98%
Table curtain (1 mm Pb): 98%
Glasses: 50 – 75%
Introduction of double dosimetry

## Slide 101
Pregnant woman radiation protection
RADIATION PROTECTION
Irradiation in utero

## Slide 102
Pregnant woman radiation protection
RADIATION PROTECTION
Patient
In the case of fluoroscopy and CT it is necessary to check whether the patient is pregnant. If the patient is pregnant, the radiologist must justify the exposition.
Wearing an apron can hinder the smooth running of the exam
The same protocols must be applied in pregnant women as in the general population.
The examination can be carried out without risk. The maximum dose is 0.2mGy. However, it is necessary to inform the patient and obtain her consent.
If the abdomen and pelvis are not directly irradiated

## Slide 103
Pregnant woman radiation protection
RADIATION PROTECTION
Patient
If the abdomen and pelvis are directly irradiated
The dose to the uterus must be evaluated a posteriori by a medical physicist and recorded by the latter in the patient's file.
Information and consent
Order of magnitude in conventional radiology: 1 to 5 mGy
Order of magnitude in fluoroscopy: About 10 mGy (embolization max 100 mGy)
Order of magnitude in CT: 5 to 50 mGy

## Slide 104
Pregnant woman radiation protection
RADIATION PROTECTION
Worker

## Slide 105
Pregnant woman radiation protection
RADIATION PROTECTION
Worker
Wearing an individual dosimeter is always compulsory. The dosimeter must be placed at the level of the abdomen (normally on the thorax) and under the lead apron*.
*No need to notify the dosimetry service

## Slide 106
CONCLUSION
Conclusion
Fluoroscopy is a challenge for radiation protection
  Highest doses delivered (patient and staff)
Interventional radiology/cardiology should be carefully monitored
  Dose reduction can be obtain most of the time (efficient protection devices)
Dose indicators allow a good control on patient exposure
All other sectors where fluoroscopy is used should be closely monitored (operating rooms, gastroenterology, urology, orthopedics…)

## Slide 107
SUMMARY
Summary

## Slide 108
SUMMARY
Summary

## Slide 109
EXERCICES
Exercices
Exercice: Examination in the abdominal area at 15 i/s
  80 kV – 200 mA – 1 ms per image (scopy)
  Distance focal spot-skin: 70 cm
  Dose at skin level?
DE [mGy]= 0.15 x (kV /100)2 x mAs x (1/SSD)2 . BSF
SSD = source skin distance [m]
BSF = back scatter factor = retrodiffusion factor (1.1 for children; 1.3 for adultes (60 – 80 kV); 1.5 for adulte (120 – 140 kV)

## Slide 110
EXERCICES
Exercices
Any idea on the dose rate to the skin:
abdominal region of a 75 kg patient
adequate geometry
in fluoroscopy mode at 15 fps
no additional copper filtration?
< 5 mGy/min
5 to 25 mGy/min
25 to 80 mGy/min
>100 mGy/min

## Slide 111
EXERCICES
Exercices
From what level of dose absorbed to the skin can the onset of erythema be predicted?
500 mGy
1000 mGy
2000 mGy
4000 mGy

## Slide 112
FLUOROSCOPY
Fluoroscopy
Any questions ?